## Description
3 Plate replicates
1st column in each plate pipetted using electronic pipette (max expected error : 2%. Avg expected error: 0.5%)
Greiner uclear black plate used for reduced autofluorescence, reduced cross fluorescence and increased sensitivity to K2Cr2O7 absorbance at 350nm
Assume electronic pipette has full accuracy and calibrate fixed tips against them. Normalize absorbance readings from electronically pipetted volumes - absorbance readings will span 0 to 1. This is to reduce variability across reads/plates
MidVolume | Stock: 1 g/L K2Cr2O7 |  |  |  | Total Stock Volume Prepared: 50 mL
Stock Volume/Water Volume (uL)
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12
A | 50/150 | ___/150 | ___/160 | ___/170 | ___/175 | ___/180 | ___/185 | ___/187.5 | ___/190
B | 40/160 | ___/150 | ___/160 | ___/170 | ___/175 | ___/180 | ___/185 | ___/187.5 | ___/190
C | 30/170 | ___/150 | ___/160 | ___/170 | ___/175 | ___/180 | ___/185 | ___/187.5 | ___/190
D | 25/175 | ___/150 | ___/160 | ___/170 | ___/175 | ___/180 | ___/185 | ___/187.5 | ___/190
E | 20/180 | ___/150 | ___/160 | ___/170 | ___/175 | ___/180 | ___/185 | ___/187.5 | ___/190
F | 15/185 | ___/150 | ___/160 | ___/170 | ___/175 | ___/180 | ___/185 | ___/187.5 | ___/190
G | 12.5/187.5 | ___/150 | ___/160 | ___/170 | ___/175 | ___/180 | ___/185 | ___/187.5 | ___/190
H | 10/190 | ___/150 | ___/160 | ___/170 | ___/175 | ___/180 | ___/185 | ___/187.5 | ___/190

## Setup
Manual
50 | 50 | 40 | 30 | 25 | 20 | 15 | 12.5 | 10
40 | 50 | 40 | 30 | 25 | 20 | 15 | 12.5 | 10
30 | 50 | 40 | 30 | 25 | 20 | 15 | 12.5 | 10
25 | 50 | 40 | 30 | 25 | 20 | 15 | 12.5 | 10
20 | 50 | 40 | 30 | 25 | 20 | 15 | 12.5 | 10
15 | 50 | 40 | 30 | 25 | 20 | 15 | 12.5 | 10
12.5 | 50 | 40 | 30 | 25 | 20 | 15 | 12.5 | 10
10 | 50 | 40 | 30 | 25 | 20 | 15 | 12.5 | 10

## Rep1
Application: Tecan i-control |  |  |  | Tecan i-control , 2.0.10.0
Device: infinite 200Pro |  |  |  | Serial number: 1903006852 |  |  |  | Serial number of connected stacker:
Firmware: V_5.31_04/18_InfiniteRX (Mar 28 2018/16.41.26) |  |  |  | MAI, V_5.31_04/18_InfiniteRX (Mar 28 2018/16.41.26)
Date: | 2020-10-02
Time: | 6:24:19 PM
System |  |  |  | TECAN-PC
User |  |  |  | TECAN-PC\tecan-user
Plate |  |  |  | Greiner 96 Flat Bottom Black Polystyrene Cat. No.: 655090/655096/655097 [GRE96fb_µClear.pdfx]
Plate-ID (Stacker)
Wait (Time) |  |  |  | 0:01
Label: rep1
Mode |  |  |  | Absorbance
Measurement Wavelength |  |  |  | 350 | nm
Bandwidth |  |  |  | 9 | nm
Number of Flashes |  |  |  | 5
Settle Time |  |  |  | 0 | ms
Part of Plate |  |  |  | A1-H10
Start Time: | 10/2/2020 6:25:19 PM
Temperature: 25 °C
<> | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
A | 1.4282 | 1.3743 | 1.1225 | 0.846 | 0.7286 | 0.6029 | 0.4718 | 0.4092 | 0.3425 | 0.0754
B | 1.1435 | 1.3686 | 1.1076 | 0.852 | 0.7225 | 0.5806 | 0.4606 | 0.393 | 0.3231 | 0.0755
C | 0.8582 | 1.3938 | 1.1362 | 0.8728 | 0.7418 | 0.5947 | 0.464 | 0.3897 | 0.3267 | 0.0756
D | 0.7317 | 1.3651 | 1.108 | 0.8454 | 0.7205 | 0.5818 | 0.4475 | 0.3897 | 0.3226 | 0.0754
E | 0.6128 | 1.3719 | 1.0997 | 0.8499 | 0.7127 | 0.5805 | 0.4505 | 0.3906 | 0.3238 | 0.0762
F | 0.4813 | 1.3652 | 1.1013 | 0.8464 | 0.7117 | 0.5841 | 0.4446 | 0.3961 | 0.3212 | 0.0762
G | 0.4132 | 1.3684 | 1.1176 | 0.8428 | 0.703 | 0.562 | 0.4416 | 0.3815 | 0.3149 | 0.0765
H | 0.3463 | 1.3724 | 1.1129 | 0.864 | 0.7228 | 0.5971 | 0.4578 | 0.3931 | 0.3323 | 0.0764
End Time: | 10/2/2020 6:25:47 PM
Wait (Time) |  |  |  | 0:00
Label: rep2
Mode |  |  |  | Absorbance
Measurement Wavelength |  |  |  | 350 | nm
Bandwidth |  |  |  | 9 | nm
Number of Flashes |  |  |  | 5
Settle Time |  |  |  | 0 | ms
Part of Plate |  |  |  | A1-H10
Start Time: | 10/2/2020 6:26:20 PM
Temperature: 25 °C
<> | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
A | 1.4319 | 1.3778 | 1.1238 | 0.8472 | 0.7309 | 0.604 | 0.4725 | 0.4099 | 0.3428 | 0.0753
B | 1.1448 | 1.3713 | 1.1093 | 0.8522 | 0.724 | 0.582 | 0.4612 | 0.3931 | 0.3235 | 0.0757
C | 0.8596 | 1.397 | 1.1389 | 0.8733 | 0.7436 | 0.5956 | 0.4646 | 0.3897 | 0.327 | 0.0755
D | 0.7318 | 1.3649 | 1.1108 | 0.8478 | 0.7208 | 0.5829 | 0.448 | 0.3899 | 0.3228 | 0.0753
E | 0.6133 | 1.3727 | 1.1011 | 0.8504 | 0.7126 | 0.5828 | 0.4506 | 0.3895 | 0.3239 | 0.0766
F | 0.4806 | 1.3668 | 1.1019 | 0.8466 | 0.7125 | 0.5848 | 0.4449 | 0.3947 | 0.3214 | 0.075
G | 0.4131 | 1.3703 | 1.1194 | 0.8457 | 0.7036 | 0.5619 | 0.4411 | 0.3814 | 0.3155 | 0.0763
H | 0.3468 | 1.3736 | 1.1121 | 0.8638 | 0.7243 | 0.5981 | 0.4583 | 0.3937 | 0.3322 | 0.0761
End Time: | 10/2/2020 6:26:46 PM
Wait (Time) |  |  |  | 0:00
Label: rep3
Mode |  |  |  | Absorbance
Measurement Wavelength |  |  |  | 350 | nm
Bandwidth |  |  |  | 9 | nm
Number of Flashes |  |  |  | 5
Settle Time |  |  |  | 0 | ms
Part of Plate |  |  |  | A1-H10
Start Time: | 10/2/2020 6:27:20 PM
Temperature: 25 °C
<> | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
A | 1.4329 | 1.3789 | 1.1249 | 0.8502 | 0.7314 | 0.6044 | 0.473 | 0.4105 | 0.3433 | 0.0753
B | 1.1494 | 1.3738 | 1.1112 | 0.8547 | 0.7252 | 0.5822 | 0.4617 | 0.3932 | 0.324 | 0.0764
C | 0.8598 | 1.3981 | 1.139 | 0.8731 | 0.744 | 0.5962 | 0.4646 | 0.3887 | 0.3274 | 0.0755
D | 0.7336 | 1.37 | 1.1121 | 0.846 | 0.7218 | 0.5831 | 0.4481 | 0.3903 | 0.3227 | 0.0753
E | 0.6137 | 1.3755 | 1.1032 | 0.8508 | 0.7134 | 0.5833 | 0.451 | 0.3898 | 0.324 | 0.0768
F | 0.4804 | 1.3649 | 1.1039 | 0.8468 | 0.7121 | 0.585 | 0.445 | 0.3944 | 0.3213 | 0.0751
G | 0.4134 | 1.368 | 1.1188 | 0.845 | 0.7043 | 0.5619 | 0.4412 | 0.3817 | 0.3158 | 0.0749
H | 0.3467 | 1.3738 | 1.1129 | 0.8643 | 0.7235 | 0.5973 | 0.4579 | 0.3924 | 0.3316 | 0.0759
End Time: | 10/2/2020 6:27:46 PM
Note on E1: Tecan.At.Common, 3.9.1.0 Tecan.At.Common.DocumentManagement, 3.9.1.0 Tecan.At.Common.DocumentManagement.Reader, 3.6.1.0 Tecan.At.Common.MCS, 3.9.1.0 Tecan.At.Common.Results, 3.9.1.0 Tecan.At.Common.UI, 3.9.1.0 Tecan.At.Communication.Common, 3.9.1.0 Tecan.At.Communication.Port.IP, 3.9.1.0 Tecan.At.Communication.Port.RS232, 3.9.1.0 Tecan.At.Communication.Port.SIM.Common, 3.9.1.0 Tecan.At.Communication.Port.USB, 3.9.1.0 Tecan.At.Communication.Server, 3.9.1.0 Tecan.At.Communication.SIM.AMR, 3.6.1.0 Tecan.At.Communication.SIM.AMRPlus, 3.6.1.0 Tecan.At.Communication.SIM.Connect, 3.9.1.0 Tecan.At.Communication.SIM.GeniosUltra, 3.6.1.0 Tecan.At.Communication.SIM.Safire3, 3.6.1.0 Tecan.At.Communication.SIM.Safire3Pro, 3.6.1.0 Tecan.At.Communication.SIM.SunriseMini, 3.6.1.0 Tecan.At.Instrument.Common, 3.9.1.0 Tecan.At.Instrument.Common.GCM, 3.9.1.0 Tecan.At.Instrument.Common.Reader, 3.6.1.0 Tecan.At.Instrument.Common.Stacker, 3.9.1.0 Tecan.At.Instrument.Gas.GCM, 3.9.1.0 Tecan.At.Instrument.GCM.Server, 3.9.1.0 Tecan.At.Instrument.Reader.AMR, 3.6.1.0 Tecan.At.Instrument.Reader.AMRPlus, 3.6.1.0 Tecan.At.Instrument.Reader.GeniosUltra, 3.6.1.0 Tecan.At.Instrument.Reader.Safire3, 3.6.1.0 Tecan.At.Instrument.Reader.Safire3Pro, 3.6.1.0 Tecan.At.Instrument.Reader.SunriseMini, 3.6.1.0 Tecan.At.Instrument.Server, 3.9.1.0 Tecan.At.Instrument.Stacker.Connect, 3.9.1.0 Tecan.At.Instrument.Stacker.Server, 3.9.1.0 Tecan.At.Measurement.BuiltInTest.Common, 3.6.1.0 Tecan.At.Measurement.Common, 3.6.1.0 Tecan.At.Measurement.Server, 3.6.1.0 Tecan.At.XFluor, 2.0.10.0 Tecan.At.XFluor.Connect.Reader, 2.0.10.0 Tecan.At.XFluor.Core, 2.0.10.0 Tecan.At.XFluor.Device, 2.0.10.0 Tecan.At.XFluor.Device.AMR, 2.0.10.0 Tecan.At.XFluor.Device.AMRPlus, 2.0.10.0 Tecan.At.XFluor.Device.GeniosUltra, 2.0.10.0 Tecan.At.XFluor.Device.Reader, 2.0.10.0 Tecan.At.XFluor.Device.Safire3, 2.0.10.0 Tecan.At.XFluor.Device.Safire3Pro, 2.0.10.0 Tecan.At.XFluor.Device.SunriseMini, 2.0.10.0 Tecan.At.XFluor.ExcelOutput, 2.0.10.0 Tecan.At.XFluor.NanoQuant, 2.0.10.0 Tecan.At.XFluor.ReaderEditor, 2.0.10.0
Note on E3: EHC, V_5.31_04/18_InfiniteRX (Mar 28 2018/16.41.26) MTP, V_5.31_04/18_InfiniteRX (Mar 28 2018/16.41.26) HCP, V_5.31_04/18_InfiniteRX (Mar 28 2018/16.41.26) LUM, V_2.20_02/2015_LUMINESCENCE (Feb 24 2015/16.09.02) MEM, V_3.00_09/11_MCR (Sep 27 2011/15.05.45) MEX, V_3.00_09/11_MCR (Sep 27 2011/15.05.10) ZSCAN, V_5.31_04/18_InfiniteRX (Mar 28 2018/16.41.26)

## Rep2
Application: Tecan i-control |  |  |  | Tecan i-control , 2.0.10.0
Device: infinite 200Pro |  |  |  | Serial number: 1903006852 |  |  |  | Serial number of connected stacker:
Firmware: V_5.31_04/18_InfiniteRX (Mar 28 2018/16.41.26) |  |  |  | MAI, V_5.31_04/18_InfiniteRX (Mar 28 2018/16.41.26)
Date: | 2020-10-02
Time: | 6:30:36 PM
System |  |  |  | TECAN-PC
User |  |  |  | TECAN-PC\tecan-user
Plate |  |  |  | Greiner 96 Flat Bottom Black Polystyrene Cat. No.: 655090/655096/655097 [GRE96fb_µClear.pdfx]
Plate-ID (Stacker)
Wait (Time) |  |  |  | 0:01
Label: rep1
Mode |  |  |  | Absorbance
Measurement Wavelength |  |  |  | 350 | nm
Bandwidth |  |  |  | 9 | nm
Number of Flashes |  |  |  | 5
Settle Time |  |  |  | 0 | ms
Part of Plate |  |  |  | A1-H10
Start Time: | 10/2/2020 6:31:36 PM
Temperature: 24.8 °C
<> | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
A | 1.4323 | 1.3864 | 1.1316 | 0.8663 | 0.7333 | 0.609 | 0.4629 | 0.3983 | 0.3432 | 0.0747
B | 1.151 | 1.3642 | 1.1148 | 0.8586 | 0.7139 | 0.5844 | 0.4557 | 0.3922 | 0.322 | 0.0756
C | 0.8572 | 1.3939 | 1.1328 | 0.8787 | 0.7295 | 0.5868 | 0.4698 | 0.3938 | 0.3369 | 0.0749
D | 0.7314 | 1.3772 | 1.114 | 0.8575 | 0.7181 | 0.5854 | 0.451 | 0.3941 | 0.3214 | 0.0755
E | 0.6098 | 1.3639 | 1.1081 | 0.8505 | 0.7166 | 0.5811 | 0.4505 | 0.3989 | 0.3212 | 0.0748
F | 0.4798 | 1.3864 | 1.1191 | 0.8542 | 0.7201 | 0.5857 | 0.4659 | 0.3928 | 0.3352 | 0.0745
G | 0.4106 | 1.3835 | 1.1251 | 0.842 | 0.7039 | 0.5705 | 0.4454 | 0.3831 | 0.3178 | 0.0754
H | 0.3475 | 1.397 | 1.1398 | 0.8664 | 0.7211 | 0.6066 | 0.4584 | 0.4012 | 0.3352 | 0.0748
End Time: | 10/2/2020 6:32:03 PM
Wait (Time) |  |  |  | 0:00
Label: rep2
Mode |  |  |  | Absorbance
Measurement Wavelength |  |  |  | 350 | nm
Bandwidth |  |  |  | 9 | nm
Number of Flashes |  |  |  | 5
Settle Time |  |  |  | 0 | ms
Part of Plate |  |  |  | A1-H10
Start Time: | 10/2/2020 6:32:36 PM
Temperature: 24.9 °C
<> | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
A | 1.4365 | 1.391 | 1.1333 | 0.87 | 0.7354 | 0.6096 | 0.462 | 0.4009 | 0.3442 | 0.0747
B | 1.153 | 1.3708 | 1.1176 | 0.8606 | 0.7153 | 0.5852 | 0.4561 | 0.3912 | 0.3224 | 0.0762
C | 0.8586 | 1.3976 | 1.1351 | 0.879 | 0.7298 | 0.5874 | 0.4716 | 0.394 | 0.3371 | 0.0749
D | 0.732 | 1.379 | 1.1149 | 0.8573 | 0.7181 | 0.5861 | 0.453 | 0.3945 | 0.3216 | 0.0745
E | 0.6105 | 1.3639 | 1.1092 | 0.8508 | 0.7178 | 0.5832 | 0.4507 | 0.3992 | 0.3223 | 0.0745
F | 0.4805 | 1.3862 | 1.1183 | 0.8548 | 0.7169 | 0.5859 | 0.465 | 0.3927 | 0.3348 | 0.0742
G | 0.4107 | 1.3808 | 1.1229 | 0.8432 | 0.7053 | 0.57 | 0.4434 | 0.3828 | 0.3161 | 0.0749
H | 0.3476 | 1.3985 | 1.1408 | 0.8663 | 0.7215 | 0.6065 | 0.4588 | 0.4007 | 0.3344 | 0.0746
End Time: | 10/2/2020 6:33:02 PM
Wait (Time) |  |  |  | 0:00
Label: rep3
Mode |  |  |  | Absorbance
Measurement Wavelength |  |  |  | 350 | nm
Bandwidth |  |  |  | 9 | nm
Number of Flashes |  |  |  | 5
Settle Time |  |  |  | 0 | ms
Part of Plate |  |  |  | A1-H10
Start Time: | 10/2/2020 6:33:36 PM
Temperature: 24.9 °C
<> | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
A | 1.4375 | 1.3896 | 1.1346 | 0.8701 | 0.7335 | 0.6099 | 0.4622 | 0.3993 | 0.3453 | 0.0745
B | 1.1536 | 1.3711 | 1.1184 | 0.862 | 0.715 | 0.5856 | 0.4558 | 0.3911 | 0.3219 | 0.0752
C | 0.8605 | 1.3979 | 1.1349 | 0.8799 | 0.7303 | 0.5882 | 0.4704 | 0.3943 | 0.3388 | 0.0749
D | 0.7332 | 1.3807 | 1.116 | 0.8585 | 0.7183 | 0.5859 | 0.4512 | 0.3944 | 0.3214 | 0.0744
E | 0.612 | 1.3646 | 1.1096 | 0.8514 | 0.7175 | 0.5822 | 0.451 | 0.3995 | 0.3206 | 0.0747
F | 0.4798 | 1.3866 | 1.12 | 0.855 | 0.7176 | 0.587 | 0.465 | 0.3931 | 0.3355 | 0.074
G | 0.4112 | 1.3847 | 1.1223 | 0.8447 | 0.7051 | 0.5708 | 0.4443 | 0.3817 | 0.3152 | 0.0738
H | 0.3482 | 1.3981 | 1.1402 | 0.8664 | 0.7209 | 0.6063 | 0.4588 | 0.4012 | 0.3335 | 0.0743
End Time: | 10/2/2020 6:34:02 PM
Note on E1: Tecan.At.Common, 3.9.1.0 Tecan.At.Common.DocumentManagement, 3.9.1.0 Tecan.At.Common.DocumentManagement.Reader, 3.6.1.0 Tecan.At.Common.MCS, 3.9.1.0 Tecan.At.Common.Results, 3.9.1.0 Tecan.At.Common.UI, 3.9.1.0 Tecan.At.Communication.Common, 3.9.1.0 Tecan.At.Communication.Port.IP, 3.9.1.0 Tecan.At.Communication.Port.RS232, 3.9.1.0 Tecan.At.Communication.Port.SIM.Common, 3.9.1.0 Tecan.At.Communication.Port.USB, 3.9.1.0 Tecan.At.Communication.Server, 3.9.1.0 Tecan.At.Communication.SIM.AMR, 3.6.1.0 Tecan.At.Communication.SIM.AMRPlus, 3.6.1.0 Tecan.At.Communication.SIM.Connect, 3.9.1.0 Tecan.At.Communication.SIM.GeniosUltra, 3.6.1.0 Tecan.At.Communication.SIM.Safire3, 3.6.1.0 Tecan.At.Communication.SIM.Safire3Pro, 3.6.1.0 Tecan.At.Communication.SIM.SunriseMini, 3.6.1.0 Tecan.At.Instrument.Common, 3.9.1.0 Tecan.At.Instrument.Common.GCM, 3.9.1.0 Tecan.At.Instrument.Common.Reader, 3.6.1.0 Tecan.At.Instrument.Common.Stacker, 3.9.1.0 Tecan.At.Instrument.Gas.GCM, 3.9.1.0 Tecan.At.Instrument.GCM.Server, 3.9.1.0 Tecan.At.Instrument.Reader.AMR, 3.6.1.0 Tecan.At.Instrument.Reader.AMRPlus, 3.6.1.0 Tecan.At.Instrument.Reader.GeniosUltra, 3.6.1.0 Tecan.At.Instrument.Reader.Safire3, 3.6.1.0 Tecan.At.Instrument.Reader.Safire3Pro, 3.6.1.0 Tecan.At.Instrument.Reader.SunriseMini, 3.6.1.0 Tecan.At.Instrument.Server, 3.9.1.0 Tecan.At.Instrument.Stacker.Connect, 3.9.1.0 Tecan.At.Instrument.Stacker.Server, 3.9.1.0 Tecan.At.Measurement.BuiltInTest.Common, 3.6.1.0 Tecan.At.Measurement.Common, 3.6.1.0 Tecan.At.Measurement.Server, 3.6.1.0 Tecan.At.XFluor, 2.0.10.0 Tecan.At.XFluor.Connect.Reader, 2.0.10.0 Tecan.At.XFluor.Core, 2.0.10.0 Tecan.At.XFluor.Device, 2.0.10.0 Tecan.At.XFluor.Device.AMR, 2.0.10.0 Tecan.At.XFluor.Device.AMRPlus, 2.0.10.0 Tecan.At.XFluor.Device.GeniosUltra, 2.0.10.0 Tecan.At.XFluor.Device.Reader, 2.0.10.0 Tecan.At.XFluor.Device.Safire3, 2.0.10.0 Tecan.At.XFluor.Device.Safire3Pro, 2.0.10.0 Tecan.At.XFluor.Device.SunriseMini, 2.0.10.0 Tecan.At.XFluor.ExcelOutput, 2.0.10.0 Tecan.At.XFluor.NanoQuant, 2.0.10.0 Tecan.At.XFluor.ReaderEditor, 2.0.10.0
Note on E3: EHC, V_5.31_04/18_InfiniteRX (Mar 28 2018/16.41.26) MTP, V_5.31_04/18_InfiniteRX (Mar 28 2018/16.41.26) HCP, V_5.31_04/18_InfiniteRX (Mar 28 2018/16.41.26) LUM, V_2.20_02/2015_LUMINESCENCE (Feb 24 2015/16.09.02) MEM, V_3.00_09/11_MCR (Sep 27 2011/15.05.45) MEX, V_3.00_09/11_MCR (Sep 27 2011/15.05.10) ZSCAN, V_5.31_04/18_InfiniteRX (Mar 28 2018/16.41.26)

## Rep3
Application: Tecan i-control |  |  |  | Tecan i-control , 2.0.10.0
Device: infinite 200Pro |  |  |  | Serial number: 1903006852 |  |  |  | Serial number of connected stacker:
Firmware: V_5.31_04/18_InfiniteRX (Mar 28 2018/16.41.26) |  |  |  | MAI, V_5.31_04/18_InfiniteRX (Mar 28 2018/16.41.26)
Date: | 2020-10-02
Time: | 6:36:52 PM
System |  |  |  | TECAN-PC
User |  |  |  | TECAN-PC\tecan-user
Plate |  |  |  | Greiner 96 Flat Bottom Black Polystyrene Cat. No.: 655090/655096/655097 [GRE96fb_µClear.pdfx]
Plate-ID (Stacker)
Wait (Time) |  |  |  | 0:01
Label: rep1
Mode |  |  |  | Absorbance
Measurement Wavelength |  |  |  | 350 | nm
Bandwidth |  |  |  | 9 | nm
Number of Flashes |  |  |  | 5
Settle Time |  |  |  | 0 | ms
Part of Plate |  |  |  | A1-H10
Start Time: | 10/2/2020 6:37:52 PM
Temperature: 24.7 °C
<> | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
A | 1.4323 | 1.3993 | 1.1344 | 0.8687 | 0.7452 | 0.6152 | 0.4709 | 0.4115 | 0.3388 | 0.0767
B | 1.158 | 1.383 | 1.1263 | 0.8641 | 0.7249 | 0.603 | 0.4638 | 0.3993 | 0.3337 | 0.0763
C | 0.8681 | 1.4096 | 1.1432 | 0.8799 | 0.7391 | 0.6169 | 0.4704 | 0.4099 | 0.3321 | 0.0775
D | 0.7382 | 1.39 | 1.1288 | 0.8669 | 0.7301 | 0.6143 | 0.4733 | 0.4072 | 0.3243 | 0.0784
E | 0.6081 | 1.3763 | 1.1138 | 0.8603 | 0.735 | 0.5991 | 0.4763 | 0.3979 | 0.3275 | 0.0773
F | 0.4843 | 1.4056 | 1.1162 | 0.8634 | 0.7431 | 0.6093 | 0.4654 | 0.4077 | 0.3341 | 0.0784
G | 0.4127 | 1.397 | 1.1256 | 0.8499 | 0.7258 | 0.5893 | 0.4544 | 0.3938 | 0.3245 | 0.0776
H | 0.3435 | 1.3882 | 1.1122 | 0.8695 | 0.7532 | 0.6099 | 0.469 | 0.4165 | 0.3399 | 0.0764
End Time: | 10/2/2020 6:38:19 PM
Wait (Time) |  |  |  | 0:00
Label: rep2
Mode |  |  |  | Absorbance
Measurement Wavelength |  |  |  | 350 | nm
Bandwidth |  |  |  | 9 | nm
Number of Flashes |  |  |  | 5
Settle Time |  |  |  | 0 | ms
Part of Plate |  |  |  | A1-H10
Start Time: | 10/2/2020 6:38:52 PM
Temperature: 24.8 °C
<> | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
A | 1.4365 | 1.4024 | 1.1358 | 0.8708 | 0.7459 | 0.6153 | 0.4708 | 0.411 | 0.3383 | 0.0764
B | 1.1575 | 1.3842 | 1.1265 | 0.8653 | 0.7254 | 0.6032 | 0.4646 | 0.399 | 0.3333 | 0.0763
C | 0.8692 | 1.4105 | 1.1443 | 0.8804 | 0.7374 | 0.616 | 0.4702 | 0.4084 | 0.3318 | 0.0773
D | 0.7379 | 1.3901 | 1.13 | 0.8667 | 0.7302 | 0.6141 | 0.4712 | 0.4065 | 0.3243 | 0.0766
E | 0.6081 | 1.376 | 1.1143 | 0.8599 | 0.7332 | 0.6005 | 0.4756 | 0.3976 | 0.3266 | 0.0779
F | 0.4827 | 1.4061 | 1.1177 | 0.8634 | 0.7418 | 0.6087 | 0.4624 | 0.4077 | 0.3332 | 0.0771
G | 0.4123 | 1.3959 | 1.1266 | 0.849 | 0.7241 | 0.5878 | 0.453 | 0.3927 | 0.3237 | 0.0768
H | 0.3427 | 1.3853 | 1.1122 | 0.8689 | 0.7512 | 0.6066 | 0.4684 | 0.4146 | 0.3379 | 0.076
End Time: | 10/2/2020 6:39:18 PM
Wait (Time) |  |  |  | 0:00
Label: rep3
Mode |  |  |  | Absorbance
Measurement Wavelength |  |  |  | 350 | nm
Bandwidth |  |  |  | 9 | nm
Number of Flashes |  |  |  | 5
Settle Time |  |  |  | 0 | ms
Part of Plate |  |  |  | A1-H10
Start Time: | 10/2/2020 6:39:52 PM
Temperature: 24.8 °C
<> | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
A | 1.4375 | 1.4029 | 1.1362 | 0.8714 | 0.7452 | 0.6154 | 0.4709 | 0.4111 | 0.3382 | 0.0764
B | 1.1596 | 1.3854 | 1.128 | 0.8665 | 0.7278 | 0.6062 | 0.4655 | 0.4014 | 0.3355 | 0.0763
C | 0.8696 | 1.4127 | 1.1439 | 0.8798 | 0.7372 | 0.6176 | 0.4707 | 0.4081 | 0.3333 | 0.0772
D | 0.7377 | 1.3897 | 1.1296 | 0.8662 | 0.7299 | 0.6128 | 0.473 | 0.4067 | 0.3247 | 0.0767
E | 0.6082 | 1.3774 | 1.1129 | 0.8599 | 0.7326 | 0.599 | 0.4766 | 0.3976 | 0.3272 | 0.077
F | 0.4822 | 1.4038 | 1.1172 | 0.8631 | 0.7408 | 0.6085 | 0.4618 | 0.407 | 0.3326 | 0.0772
G | 0.4124 | 1.3962 | 1.1244 | 0.8499 | 0.7234 | 0.5885 | 0.4526 | 0.3924 | 0.3235 | 0.0761
H | 0.3425 | 1.3861 | 1.112 | 0.8699 | 0.7534 | 0.607 | 0.4667 | 0.4148 | 0.3397 | 0.0761
End Time: | 10/2/2020 6:40:18 PM
Note on E1: Tecan.At.Common, 3.9.1.0 Tecan.At.Common.DocumentManagement, 3.9.1.0 Tecan.At.Common.DocumentManagement.Reader, 3.6.1.0 Tecan.At.Common.MCS, 3.9.1.0 Tecan.At.Common.Results, 3.9.1.0 Tecan.At.Common.UI, 3.9.1.0 Tecan.At.Communication.Common, 3.9.1.0 Tecan.At.Communication.Port.IP, 3.9.1.0 Tecan.At.Communication.Port.RS232, 3.9.1.0 Tecan.At.Communication.Port.SIM.Common, 3.9.1.0 Tecan.At.Communication.Port.USB, 3.9.1.0 Tecan.At.Communication.Server, 3.9.1.0 Tecan.At.Communication.SIM.AMR, 3.6.1.0 Tecan.At.Communication.SIM.AMRPlus, 3.6.1.0 Tecan.At.Communication.SIM.Connect, 3.9.1.0 Tecan.At.Communication.SIM.GeniosUltra, 3.6.1.0 Tecan.At.Communication.SIM.Safire3, 3.6.1.0 Tecan.At.Communication.SIM.Safire3Pro, 3.6.1.0 Tecan.At.Communication.SIM.SunriseMini, 3.6.1.0 Tecan.At.Instrument.Common, 3.9.1.0 Tecan.At.Instrument.Common.GCM, 3.9.1.0 Tecan.At.Instrument.Common.Reader, 3.6.1.0 Tecan.At.Instrument.Common.Stacker, 3.9.1.0 Tecan.At.Instrument.Gas.GCM, 3.9.1.0 Tecan.At.Instrument.GCM.Server, 3.9.1.0 Tecan.At.Instrument.Reader.AMR, 3.6.1.0 Tecan.At.Instrument.Reader.AMRPlus, 3.6.1.0 Tecan.At.Instrument.Reader.GeniosUltra, 3.6.1.0 Tecan.At.Instrument.Reader.Safire3, 3.6.1.0 Tecan.At.Instrument.Reader.Safire3Pro, 3.6.1.0 Tecan.At.Instrument.Reader.SunriseMini, 3.6.1.0 Tecan.At.Instrument.Server, 3.9.1.0 Tecan.At.Instrument.Stacker.Connect, 3.9.1.0 Tecan.At.Instrument.Stacker.Server, 3.9.1.0 Tecan.At.Measurement.BuiltInTest.Common, 3.6.1.0 Tecan.At.Measurement.Common, 3.6.1.0 Tecan.At.Measurement.Server, 3.6.1.0 Tecan.At.XFluor, 2.0.10.0 Tecan.At.XFluor.Connect.Reader, 2.0.10.0 Tecan.At.XFluor.Core, 2.0.10.0 Tecan.At.XFluor.Device, 2.0.10.0 Tecan.At.XFluor.Device.AMR, 2.0.10.0 Tecan.At.XFluor.Device.AMRPlus, 2.0.10.0 Tecan.At.XFluor.Device.GeniosUltra, 2.0.10.0 Tecan.At.XFluor.Device.Reader, 2.0.10.0 Tecan.At.XFluor.Device.Safire3, 2.0.10.0 Tecan.At.XFluor.Device.Safire3Pro, 2.0.10.0 Tecan.At.XFluor.Device.SunriseMini, 2.0.10.0 Tecan.At.XFluor.ExcelOutput, 2.0.10.0 Tecan.At.XFluor.NanoQuant, 2.0.10.0 Tecan.At.XFluor.ReaderEditor, 2.0.10.0
Note on E3: EHC, V_5.31_04/18_InfiniteRX (Mar 28 2018/16.41.26) MTP, V_5.31_04/18_InfiniteRX (Mar 28 2018/16.41.26) HCP, V_5.31_04/18_InfiniteRX (Mar 28 2018/16.41.26) LUM, V_2.20_02/2015_LUMINESCENCE (Feb 24 2015/16.09.02) MEM, V_3.00_09/11_MCR (Sep 27 2011/15.05.45) MEX, V_3.00_09/11_MCR (Sep 27 2011/15.05.10) ZSCAN, V_5.31_04/18_InfiniteRX (Mar 28 2018/16.41.26)