## Description
3 Plate replicates
1st column in each plate pipetted using electronic pipette (max expected error : 2%. Avg expected error: 0.5%)
Greiner uclear black plate used for reduced autofluorescence, reduced cross fluorescence and increased sensitivity to K2Cr2O7 absorbance at 350nm
Assume electronic pipette has full accuracy and calibrate fixed tips against them. Normalize absorbance readings from electronically pipetted volumes - absorbance readings will span 0 to 1. This is to reduce variability across reads/plates
MidVolume | Stock: 2 g/L K2Cr2O7 |  |  |  | Total Stock Volume Prepared: 50 mL
Stock Volume/Water Volume (uL)
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12
A | 10/190 | ___/190 | ___/191 | ___/192 | ___/193 | ___/194 | ___/195 | ___/196 | ___/197
B | 9/191 | ___/190 | ___/191 | ___/192 | ___/193 | ___/194 | ___/195 | ___/196 | ___/197
C | 8/192 | ___/190 | ___/191 | ___/192 | ___/193 | ___/194 | ___/195 | ___/196 | ___/197
D | 7/193 | ___/190 | ___/191 | ___/192 | ___/193 | ___/194 | ___/195 | ___/196 | ___/197
E | 6/194 | ___/190 | ___/191 | ___/192 | ___/193 | ___/194 | ___/195 | ___/196 | ___/197
F | 5/195 | ___/190 | ___/191 | ___/192 | ___/193 | ___/194 | ___/195 | ___/196 | ___/197
G | 4/196 | ___/190 | ___/191 | ___/192 | ___/193 | ___/194 | ___/195 | ___/196 | ___/197
H | 3/197 | ___/190 | ___/191 | ___/192 | ___/193 | ___/194 | ___/195 | ___/196 | ___/197

## Setup
Manual
10 | 10 | 9 | 8 | 7 | 6 | 5 | 4 | 3
9 | 10 | 9 | 8 | 7 | 6 | 5 | 4 | 3
8 | 10 | 9 | 8 | 7 | 6 | 5 | 4 | 3
7 | 10 | 9 | 8 | 7 | 6 | 5 | 4 | 3
6 | 10 | 9 | 8 | 7 | 6 | 5 | 4 | 3
5 | 10 | 9 | 8 | 7 | 6 | 5 | 4 | 3
4 | 10 | 9 | 8 | 7 | 6 | 5 | 4 | 3
3 | 10 | 9 | 8 | 7 | 6 | 5 | 4 | 3

## Rep1
Application: Tecan i-control |  |  |  | Tecan i-control , 2.0.10.0
Device: infinite 200Pro |  |  |  | Serial number: 1903006852 |  |  |  | Serial number of connected stacker:
Firmware: V_5.31_04/18_InfiniteRX (Mar 28 2018/16.41.26) |  |  |  | MAI, V_5.31_04/18_InfiniteRX (Mar 28 2018/16.41.26)
Date: | 2021-02-15
Time: | 5:05:32 PM
System |  |  |  | TECAN-PC
User |  |  |  | TECAN-PC\tecan-user
Plate |  |  |  | Greiner 96 Flat Bottom Black Polystyrene Cat. No.: 655090/655096/655097 [GRE96fb_µClear.pdfx]
Plate-ID (Stacker)
Wait (Time) |  |  |  | 0:01
Label: rep
Mode |  |  |  | Absorbance
Measurement Wavelength |  |  |  | 350 | nm
Bandwidth |  |  |  | 9 | nm
Number of Flashes |  |  |  | 5
Settle Time |  |  |  | 0 | ms
Part of Plate |  |  |  | A1-H9
Start Time: | 2/15/2021 5:06:32 PM
Temperature: 37.4 °C
<> | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9
A | 0.7375 | 0.7021 | 0.6344 | 0.5537 | 0.4923 | 0.4185 | 0.362 | 0.3061 | 0.2728
B | 0.6402 | 0.6742 | 0.6228 | 0.573 | 0.5018 | 0.4337 | 0.3985 | 0.3135 | 0.2602
C | 0.5899 | 0.6983 | 0.6485 | 0.5465 | 0.5025 | 0.4457 | 0.3784 | 0.3023 | 0.2527
D | 0.5472 | 0.6444 | 0.5866 | 0.5236 | 0.471 | 0.4114 | 0.3702 | 0.3059 | 0.247
E | 0.4677 | 0.6743 | 0.62 | 0.5579 | 0.4741 | 0.4275 | 0.3551 | 0.3148 | 0.2668
F | 0.4004 | 0.6789 | 0.6309 | 0.5594 | 0.5088 | 0.4302 | 0.3714 | 0.3246 | 0.2701
G | 0.341 | 0.6662 | 0.5741 | 0.5107 | 0.4749 | 0.3847 | 0.3459 | 0.2906 | 0.2309
H | 0.277 | 0.6954 | 0.6643 | 0.6037 | 0.5244 | 0.4477 | 0.394 | 0.3251 | 0.2798
End Time: | 2/15/2021 5:06:57 PM
Wait (Time) |  |  |  | 0:00
Label: rep_Copy1
Mode |  |  |  | Absorbance
Measurement Wavelength |  |  |  | 350 | nm
Bandwidth |  |  |  | 9 | nm
Number of Flashes |  |  |  | 5
Settle Time |  |  |  | 0 | ms
Part of Plate |  |  |  | A1-H9
Start Time: | 2/15/2021 5:07:30 PM
Temperature: 36.8 °C
<> | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9
A | 0.737 | 0.7017 | 0.6336 | 0.5536 | 0.4921 | 0.4176 | 0.3606 | 0.3064 | 0.2728
B | 0.6392 | 0.6745 | 0.6222 | 0.5712 | 0.5034 | 0.4328 | 0.3986 | 0.3137 | 0.2593
C | 0.5896 | 0.6983 | 0.6477 | 0.5474 | 0.5033 | 0.4451 | 0.3782 | 0.3025 | 0.2517
D | 0.5468 | 0.644 | 0.5872 | 0.5231 | 0.4706 | 0.4116 | 0.3705 | 0.3054 | 0.2478
E | 0.4672 | 0.6751 | 0.6198 | 0.5578 | 0.4727 | 0.4274 | 0.3552 | 0.3147 | 0.2655
F | 0.3994 | 0.6796 | 0.6297 | 0.5595 | 0.5096 | 0.4302 | 0.3707 | 0.325 | 0.2698
G | 0.3412 | 0.667 | 0.5741 | 0.5112 | 0.4749 | 0.3844 | 0.3458 | 0.2907 | 0.2319
H | 0.2776 | 0.6952 | 0.662 | 0.6031 | 0.5252 | 0.4479 | 0.3941 | 0.3251 | 0.2791
End Time: | 2/15/2021 5:07:55 PM
Wait (Time) |  |  |  | 0:00
Label: rep_Copy2
Mode |  |  |  | Absorbance
Measurement Wavelength |  |  |  | 350 | nm
Bandwidth |  |  |  | 9 | nm
Number of Flashes |  |  |  | 5
Settle Time |  |  |  | 0 | ms
Part of Plate |  |  |  | A1-H9
Start Time: | 2/15/2021 5:08:28 PM
Temperature: 37.6 °C
<> | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9
A | 0.7368 | 0.7006 | 0.6335 | 0.553 | 0.4918 | 0.4189 | 0.3607 | 0.3059 | 0.2746
B | 0.6381 | 0.6748 | 0.6231 | 0.5723 | 0.5022 | 0.4329 | 0.3985 | 0.3139 | 0.259
C | 0.5896 | 0.6991 | 0.6471 | 0.5464 | 0.5037 | 0.4455 | 0.3785 | 0.3027 | 0.2519
D | 0.549 | 0.6466 | 0.588 | 0.5237 | 0.471 | 0.412 | 0.3704 | 0.305 | 0.247
E | 0.4672 | 0.6751 | 0.6198 | 0.5579 | 0.4735 | 0.4284 | 0.3555 | 0.3163 | 0.2656
F | 0.3982 | 0.6793 | 0.629 | 0.5579 | 0.5066 | 0.4296 | 0.3707 | 0.3241 | 0.2698
G | 0.3411 | 0.6671 | 0.5743 | 0.5114 | 0.4746 | 0.3846 | 0.3459 | 0.2905 | 0.2304
H | 0.2774 | 0.6954 | 0.6612 | 0.6018 | 0.5239 | 0.4474 | 0.3934 | 0.3252 | 0.2794
End Time: | 2/15/2021 5:08:52 PM
Note on E1: Tecan.At.Common, 3.9.1.0 Tecan.At.Common.DocumentManagement, 3.9.1.0 Tecan.At.Common.DocumentManagement.Reader, 3.6.1.0 Tecan.At.Common.MCS, 3.9.1.0 Tecan.At.Common.Results, 3.9.1.0 Tecan.At.Common.UI, 3.9.1.0 Tecan.At.Communication.Common, 3.9.1.0 Tecan.At.Communication.Port.IP, 3.9.1.0 Tecan.At.Communication.Port.RS232, 3.9.1.0 Tecan.At.Communication.Port.SIM.Common, 3.9.1.0 Tecan.At.Communication.Port.USB, 3.9.1.0 Tecan.At.Communication.Server, 3.9.1.0 Tecan.At.Communication.SIM.AMR, 3.6.1.0 Tecan.At.Communication.SIM.AMRPlus, 3.6.1.0 Tecan.At.Communication.SIM.Connect, 3.9.1.0 Tecan.At.Communication.SIM.GeniosUltra, 3.6.1.0 Tecan.At.Communication.SIM.Safire3, 3.6.1.0 Tecan.At.Communication.SIM.Safire3Pro, 3.6.1.0 Tecan.At.Communication.SIM.SunriseMini, 3.6.1.0 Tecan.At.Instrument.Common, 3.9.1.0 Tecan.At.Instrument.Common.GCM, 3.9.1.0 Tecan.At.Instrument.Common.Reader, 3.6.1.0 Tecan.At.Instrument.Common.Stacker, 3.9.1.0 Tecan.At.Instrument.Gas.GCM, 3.9.1.0 Tecan.At.Instrument.GCM.Server, 3.9.1.0 Tecan.At.Instrument.Reader.AMR, 3.6.1.0 Tecan.At.Instrument.Reader.AMRPlus, 3.6.1.0 Tecan.At.Instrument.Reader.GeniosUltra, 3.6.1.0 Tecan.At.Instrument.Reader.Safire3, 3.6.1.0 Tecan.At.Instrument.Reader.Safire3Pro, 3.6.1.0 Tecan.At.Instrument.Reader.SunriseMini, 3.6.1.0 Tecan.At.Instrument.Server, 3.9.1.0 Tecan.At.Instrument.Stacker.Connect, 3.9.1.0 Tecan.At.Instrument.Stacker.Server, 3.9.1.0 Tecan.At.Measurement.BuiltInTest.Common, 3.6.1.0 Tecan.At.Measurement.Common, 3.6.1.0 Tecan.At.Measurement.Server, 3.6.1.0 Tecan.At.XFluor, 2.0.10.0 Tecan.At.XFluor.Connect.Reader, 2.0.10.0 Tecan.At.XFluor.Core, 2.0.10.0 Tecan.At.XFluor.Device, 2.0.10.0 Tecan.At.XFluor.Device.AMR, 2.0.10.0 Tecan.At.XFluor.Device.AMRPlus, 2.0.10.0 Tecan.At.XFluor.Device.GeniosUltra, 2.0.10.0 Tecan.At.XFluor.Device.Reader, 2.0.10.0 Tecan.At.XFluor.Device.Safire3, 2.0.10.0 Tecan.At.XFluor.Device.Safire3Pro, 2.0.10.0 Tecan.At.XFluor.Device.SunriseMini, 2.0.10.0 Tecan.At.XFluor.ExcelOutput, 2.0.10.0 Tecan.At.XFluor.NanoQuant, 2.0.10.0 Tecan.At.XFluor.ReaderEditor, 2.0.10.0
Note on E3: EHC, V_5.31_04/18_InfiniteRX (Mar 28 2018/16.41.26) MTP, V_5.31_04/18_InfiniteRX (Mar 28 2018/16.41.26) HCP, V_5.31_04/18_InfiniteRX (Mar 28 2018/16.41.26) LUM, V_2.20_02/2015_LUMINESCENCE (Feb 24 2015/16.09.02) MEM, V_3.00_09/11_MCR (Sep 27 2011/15.05.45) MEX, V_3.00_09/11_MCR (Sep 27 2011/15.05.10) ZSCAN, V_5.31_04/18_InfiniteRX (Mar 28 2018/16.41.26)

## Rep2
Application: Tecan i-control |  |  |  | Tecan i-control , 2.0.10.0
Device: infinite 200Pro |  |  |  | Serial number: 1903006852 |  |  |  | Serial number of connected stacker:
Firmware: V_5.31_04/18_InfiniteRX (Mar 28 2018/16.41.26) |  |  |  | MAI, V_5.31_04/18_InfiniteRX (Mar 28 2018/16.41.26)
Date: | 2021-02-15
Time: | 5:11:15 PM
System |  |  |  | TECAN-PC
User |  |  |  | TECAN-PC\tecan-user
Plate |  |  |  | Greiner 96 Flat Bottom Black Polystyrene Cat. No.: 655090/655096/655097 [GRE96fb_µClear.pdfx]
Plate-ID (Stacker)
Wait (Time) |  |  |  | 0:01
Label: rep
Mode |  |  |  | Absorbance
Measurement Wavelength |  |  |  | 350 | nm
Bandwidth |  |  |  | 9 | nm
Number of Flashes |  |  |  | 5
Settle Time |  |  |  | 0 | ms
Part of Plate |  |  |  | A1-H9
Start Time: | 2/15/2021 5:12:15 PM
Temperature: 37 °C
<> | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9
A | 0.7436 | 0.6763 | 0.6083 | 0.5517 | 0.5069 | 0.4133 | 0.3589 | 0.3031 | 0.2488
B | 0.6725 | 0.6958 | 0.6451 | 0.5775 | 0.4882 | 0.4347 | 0.3798 | 0.3207 | 0.263
C | 0.5816 | 0.6846 | 0.6214 | 0.564 | 0.5125 | 0.447 | 0.3872 | 0.3288 | 0.2604
D | 0.5377 | 0.6654 | 0.6036 | 0.5429 | 0.4905 | 0.4251 | 0.3703 | 0.3121 | 0.2469
E | 0.4701 | 0.6857 | 0.6021 | 0.5408 | 0.4792 | 0.4373 | 0.3524 | 0.3042 | 0.2527
F | 0.4085 | 0.695 | 0.6245 | 0.5707 | 0.5086 | 0.4365 | 0.3747 | 0.3185 | 0.267
G | 0.3418 | 0.667 | 0.5719 | 0.5149 | 0.4535 | 0.4152 | 0.3297 | 0.2825 | 0.232
H | 0.277 | 0.6883 | 0.6397 | 0.5908 | 0.5277 | 0.4774 | 0.4103 | 0.3422 | 0.2777
End Time: | 2/15/2021 5:12:39 PM
Wait (Time) |  |  |  | 0:00
Label: rep_Copy1
Mode |  |  |  | Absorbance
Measurement Wavelength |  |  |  | 350 | nm
Bandwidth |  |  |  | 9 | nm
Number of Flashes |  |  |  | 5
Settle Time |  |  |  | 0 | ms
Part of Plate |  |  |  | A1-H9
Start Time: | 2/15/2021 5:13:13 PM
Temperature: 37.5 °C
<> | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9
A | 0.7432 | 0.6767 | 0.6085 | 0.5516 | 0.5081 | 0.4132 | 0.3584 | 0.3029 | 0.2485
B | 0.672 | 0.6957 | 0.6443 | 0.5784 | 0.4877 | 0.4341 | 0.3802 | 0.3209 | 0.2613
C | 0.5809 | 0.6852 | 0.6214 | 0.5631 | 0.5123 | 0.4465 | 0.3871 | 0.3289 | 0.2603
D | 0.5382 | 0.6663 | 0.6038 | 0.5443 | 0.4906 | 0.4242 | 0.369 | 0.3114 | 0.2468
E | 0.4703 | 0.6856 | 0.6018 | 0.5401 | 0.4776 | 0.4371 | 0.3524 | 0.3045 | 0.2518
F | 0.4092 | 0.695 | 0.6251 | 0.5711 | 0.5081 | 0.4376 | 0.3753 | 0.3206 | 0.2669
G | 0.3419 | 0.6675 | 0.571 | 0.5147 | 0.4535 | 0.4153 | 0.3298 | 0.2815 | 0.2318
H | 0.2799 | 0.6866 | 0.6385 | 0.5889 | 0.529 | 0.4758 | 0.4092 | 0.3417 | 0.2775
End Time: | 2/15/2021 5:13:37 PM
Wait (Time) |  |  |  | 0:00
Label: rep_Copy2
Mode |  |  |  | Absorbance
Measurement Wavelength |  |  |  | 350 | nm
Bandwidth |  |  |  | 9 | nm
Number of Flashes |  |  |  | 5
Settle Time |  |  |  | 0 | ms
Part of Plate |  |  |  | A1-H9
Start Time: | 2/15/2021 5:14:11 PM
Temperature: 37.5 °C
<> | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9
A | 0.7435 | 0.6759 | 0.6083 | 0.5523 | 0.5065 | 0.4129 | 0.3581 | 0.3039 | 0.249
B | 0.6721 | 0.6951 | 0.6448 | 0.5783 | 0.4882 | 0.4341 | 0.3799 | 0.3208 | 0.2609
C | 0.5807 | 0.6849 | 0.6222 | 0.5636 | 0.5124 | 0.4482 | 0.3874 | 0.3284 | 0.2604
D | 0.5383 | 0.665 | 0.6036 | 0.5435 | 0.49 | 0.4244 | 0.3695 | 0.3115 | 0.2468
E | 0.4696 | 0.6862 | 0.6019 | 0.5404 | 0.4777 | 0.4372 | 0.3526 | 0.3038 | 0.2523
F | 0.4102 | 0.6962 | 0.6266 | 0.5715 | 0.5084 | 0.4376 | 0.3757 | 0.3189 | 0.2673
G | 0.3423 | 0.6677 | 0.5715 | 0.5164 | 0.4539 | 0.4151 | 0.3302 | 0.2813 | 0.2315
H | 0.2776 | 0.6871 | 0.6401 | 0.5891 | 0.5258 | 0.4756 | 0.4116 | 0.3421 | 0.2771
End Time: | 2/15/2021 5:14:35 PM
Note on E1: Tecan.At.Common, 3.9.1.0 Tecan.At.Common.DocumentManagement, 3.9.1.0 Tecan.At.Common.DocumentManagement.Reader, 3.6.1.0 Tecan.At.Common.MCS, 3.9.1.0 Tecan.At.Common.Results, 3.9.1.0 Tecan.At.Common.UI, 3.9.1.0 Tecan.At.Communication.Common, 3.9.1.0 Tecan.At.Communication.Port.IP, 3.9.1.0 Tecan.At.Communication.Port.RS232, 3.9.1.0 Tecan.At.Communication.Port.SIM.Common, 3.9.1.0 Tecan.At.Communication.Port.USB, 3.9.1.0 Tecan.At.Communication.Server, 3.9.1.0 Tecan.At.Communication.SIM.AMR, 3.6.1.0 Tecan.At.Communication.SIM.AMRPlus, 3.6.1.0 Tecan.At.Communication.SIM.Connect, 3.9.1.0 Tecan.At.Communication.SIM.GeniosUltra, 3.6.1.0 Tecan.At.Communication.SIM.Safire3, 3.6.1.0 Tecan.At.Communication.SIM.Safire3Pro, 3.6.1.0 Tecan.At.Communication.SIM.SunriseMini, 3.6.1.0 Tecan.At.Instrument.Common, 3.9.1.0 Tecan.At.Instrument.Common.GCM, 3.9.1.0 Tecan.At.Instrument.Common.Reader, 3.6.1.0 Tecan.At.Instrument.Common.Stacker, 3.9.1.0 Tecan.At.Instrument.Gas.GCM, 3.9.1.0 Tecan.At.Instrument.GCM.Server, 3.9.1.0 Tecan.At.Instrument.Reader.AMR, 3.6.1.0 Tecan.At.Instrument.Reader.AMRPlus, 3.6.1.0 Tecan.At.Instrument.Reader.GeniosUltra, 3.6.1.0 Tecan.At.Instrument.Reader.Safire3, 3.6.1.0 Tecan.At.Instrument.Reader.Safire3Pro, 3.6.1.0 Tecan.At.Instrument.Reader.SunriseMini, 3.6.1.0 Tecan.At.Instrument.Server, 3.9.1.0 Tecan.At.Instrument.Stacker.Connect, 3.9.1.0 Tecan.At.Instrument.Stacker.Server, 3.9.1.0 Tecan.At.Measurement.BuiltInTest.Common, 3.6.1.0 Tecan.At.Measurement.Common, 3.6.1.0 Tecan.At.Measurement.Server, 3.6.1.0 Tecan.At.XFluor, 2.0.10.0 Tecan.At.XFluor.Connect.Reader, 2.0.10.0 Tecan.At.XFluor.Core, 2.0.10.0 Tecan.At.XFluor.Device, 2.0.10.0 Tecan.At.XFluor.Device.AMR, 2.0.10.0 Tecan.At.XFluor.Device.AMRPlus, 2.0.10.0 Tecan.At.XFluor.Device.GeniosUltra, 2.0.10.0 Tecan.At.XFluor.Device.Reader, 2.0.10.0 Tecan.At.XFluor.Device.Safire3, 2.0.10.0 Tecan.At.XFluor.Device.Safire3Pro, 2.0.10.0 Tecan.At.XFluor.Device.SunriseMini, 2.0.10.0 Tecan.At.XFluor.ExcelOutput, 2.0.10.0 Tecan.At.XFluor.NanoQuant, 2.0.10.0 Tecan.At.XFluor.ReaderEditor, 2.0.10.0
Note on E3: EHC, V_5.31_04/18_InfiniteRX (Mar 28 2018/16.41.26) MTP, V_5.31_04/18_InfiniteRX (Mar 28 2018/16.41.26) HCP, V_5.31_04/18_InfiniteRX (Mar 28 2018/16.41.26) LUM, V_2.20_02/2015_LUMINESCENCE (Feb 24 2015/16.09.02) MEM, V_3.00_09/11_MCR (Sep 27 2011/15.05.45) MEX, V_3.00_09/11_MCR (Sep 27 2011/15.05.10) ZSCAN, V_5.31_04/18_InfiniteRX (Mar 28 2018/16.41.26)

## Rep3
Application: Tecan i-control |  |  |  | Tecan i-control , 2.0.10.0
Device: infinite 200Pro |  |  |  | Serial number: 1903006852 |  |  |  | Serial number of connected stacker:
Firmware: V_5.31_04/18_InfiniteRX (Mar 28 2018/16.41.26) |  |  |  | MAI, V_5.31_04/18_InfiniteRX (Mar 28 2018/16.41.26)
Date: | 2021-02-15
Time: | 5:16:51 PM
System |  |  |  | TECAN-PC
User |  |  |  | TECAN-PC\tecan-user
Plate |  |  |  | Greiner 96 Flat Bottom Black Polystyrene Cat. No.: 655090/655096/655097 [GRE96fb_µClear.pdfx]
Plate-ID (Stacker)
Wait (Time) |  |  |  | 0:01
Label: rep
Mode |  |  |  | Absorbance
Measurement Wavelength |  |  |  | 350 | nm
Bandwidth |  |  |  | 9 | nm
Number of Flashes |  |  |  | 5
Settle Time |  |  |  | 0 | ms
Part of Plate |  |  |  | A1-H9
Start Time: | 2/15/2021 5:17:51 PM
Temperature: 36.8 °C
<> | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9
A | 0.7512 | 0.705 | 0.6198 | 0.5386 | 0.4928 | 0.4291 | 0.3537 | 0.3003 | 0.2477
B | 0.6845 | 0.6706 | 0.6128 | 0.5541 | 0.5102 | 0.437 | 0.3721 | 0.3308 | 0.2573
C | 0.599 | 0.6962 | 0.6194 | 0.5531 | 0.4917 | 0.4327 | 0.3887 | 0.3114 | 0.258
D | 0.5436 | 0.6552 | 0.5979 | 0.5488 | 0.4813 | 0.4394 | 0.3533 | 0.2906 | 0.2562
E | 0.4721 | 0.6561 | 0.5989 | 0.5426 | 0.4662 | 0.4259 | 0.3545 | 0.3002 | 0.2545
F | 0.4079 | 0.6776 | 0.6219 | 0.5657 | 0.5059 | 0.4307 | 0.3762 | 0.312 | 0.2551
G | 0.3481 | 0.6442 | 0.6139 | 0.5342 | 0.4823 | 0.3846 | 0.3426 | 0.2746 | 0.2321
H | 0.2792 | 0.7189 | 0.6208 | 0.5734 | 0.5198 | 0.4895 | 0.4089 | 0.3587 | 0.2879
End Time: | 2/15/2021 5:18:16 PM
Wait (Time) |  |  |  | 0:00
Label: rep_Copy1
Mode |  |  |  | Absorbance
Measurement Wavelength |  |  |  | 350 | nm
Bandwidth |  |  |  | 9 | nm
Number of Flashes |  |  |  | 5
Settle Time |  |  |  | 0 | ms
Part of Plate |  |  |  | A1-H9
Start Time: | 2/15/2021 5:18:49 PM
Temperature: 37.7 °C
<> | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9
A | 0.7499 | 0.7028 | 0.6197 | 0.5379 | 0.4919 | 0.4286 | 0.3534 | 0.3001 | 0.2472
B | 0.684 | 0.6714 | 0.6129 | 0.5561 | 0.5104 | 0.4376 | 0.3722 | 0.3307 | 0.2563
C | 0.5983 | 0.6974 | 0.619 | 0.5516 | 0.4919 | 0.432 | 0.3869 | 0.3112 | 0.2574
D | 0.5431 | 0.6559 | 0.5964 | 0.5486 | 0.4814 | 0.441 | 0.3531 | 0.2905 | 0.2551
E | 0.4716 | 0.6554 | 0.5988 | 0.5439 | 0.4659 | 0.4267 | 0.3534 | 0.2996 | 0.2533
F | 0.4073 | 0.6774 | 0.6204 | 0.5661 | 0.504 | 0.4327 | 0.3761 | 0.3118 | 0.254
G | 0.3483 | 0.6434 | 0.6145 | 0.5339 | 0.4823 | 0.3861 | 0.3425 | 0.2747 | 0.2311
H | 0.2799 | 0.7201 | 0.621 | 0.5727 | 0.5183 | 0.4891 | 0.4084 | 0.3588 | 0.2878
End Time: | 2/15/2021 5:19:14 PM
Wait (Time) |  |  |  | 0:00
Label: rep_Copy2
Mode |  |  |  | Absorbance
Measurement Wavelength |  |  |  | 350 | nm
Bandwidth |  |  |  | 9 | nm
Number of Flashes |  |  |  | 5
Settle Time |  |  |  | 0 | ms
Part of Plate |  |  |  | A1-H9
Start Time: | 2/15/2021 5:19:47 PM
Temperature: 37.2 °C
<> | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9
A | 0.7504 | 0.7032 | 0.6186 | 0.5392 | 0.4922 | 0.4282 | 0.3531 | 0.2999 | 0.2471
B | 0.6836 | 0.6711 | 0.6139 | 0.5548 | 0.5107 | 0.4382 | 0.3723 | 0.331 | 0.2561
C | 0.5981 | 0.696 | 0.6188 | 0.5528 | 0.4922 | 0.431 | 0.3867 | 0.3112 | 0.2576
D | 0.5429 | 0.6551 | 0.5983 | 0.5482 | 0.4808 | 0.4392 | 0.3528 | 0.2904 | 0.2546
E | 0.471 | 0.6574 | 0.5984 | 0.5427 | 0.4666 | 0.4269 | 0.3535 | 0.2999 | 0.2536
F | 0.4078 | 0.6775 | 0.6208 | 0.5657 | 0.5037 | 0.4304 | 0.3756 | 0.3115 | 0.2539
G | 0.3472 | 0.6428 | 0.6137 | 0.5342 | 0.4832 | 0.3851 | 0.3437 | 0.2746 | 0.2309
H | 0.2793 | 0.7191 | 0.6195 | 0.5716 | 0.5178 | 0.4895 | 0.4072 | 0.3582 | 0.2877
End Time: | 2/15/2021 5:20:11 PM
Note on E1: Tecan.At.Common, 3.9.1.0 Tecan.At.Common.DocumentManagement, 3.9.1.0 Tecan.At.Common.DocumentManagement.Reader, 3.6.1.0 Tecan.At.Common.MCS, 3.9.1.0 Tecan.At.Common.Results, 3.9.1.0 Tecan.At.Common.UI, 3.9.1.0 Tecan.At.Communication.Common, 3.9.1.0 Tecan.At.Communication.Port.IP, 3.9.1.0 Tecan.At.Communication.Port.RS232, 3.9.1.0 Tecan.At.Communication.Port.SIM.Common, 3.9.1.0 Tecan.At.Communication.Port.USB, 3.9.1.0 Tecan.At.Communication.Server, 3.9.1.0 Tecan.At.Communication.SIM.AMR, 3.6.1.0 Tecan.At.Communication.SIM.AMRPlus, 3.6.1.0 Tecan.At.Communication.SIM.Connect, 3.9.1.0 Tecan.At.Communication.SIM.GeniosUltra, 3.6.1.0 Tecan.At.Communication.SIM.Safire3, 3.6.1.0 Tecan.At.Communication.SIM.Safire3Pro, 3.6.1.0 Tecan.At.Communication.SIM.SunriseMini, 3.6.1.0 Tecan.At.Instrument.Common, 3.9.1.0 Tecan.At.Instrument.Common.GCM, 3.9.1.0 Tecan.At.Instrument.Common.Reader, 3.6.1.0 Tecan.At.Instrument.Common.Stacker, 3.9.1.0 Tecan.At.Instrument.Gas.GCM, 3.9.1.0 Tecan.At.Instrument.GCM.Server, 3.9.1.0 Tecan.At.Instrument.Reader.AMR, 3.6.1.0 Tecan.At.Instrument.Reader.AMRPlus, 3.6.1.0 Tecan.At.Instrument.Reader.GeniosUltra, 3.6.1.0 Tecan.At.Instrument.Reader.Safire3, 3.6.1.0 Tecan.At.Instrument.Reader.Safire3Pro, 3.6.1.0 Tecan.At.Instrument.Reader.SunriseMini, 3.6.1.0 Tecan.At.Instrument.Server, 3.9.1.0 Tecan.At.Instrument.Stacker.Connect, 3.9.1.0 Tecan.At.Instrument.Stacker.Server, 3.9.1.0 Tecan.At.Measurement.BuiltInTest.Common, 3.6.1.0 Tecan.At.Measurement.Common, 3.6.1.0 Tecan.At.Measurement.Server, 3.6.1.0 Tecan.At.XFluor, 2.0.10.0 Tecan.At.XFluor.Connect.Reader, 2.0.10.0 Tecan.At.XFluor.Core, 2.0.10.0 Tecan.At.XFluor.Device, 2.0.10.0 Tecan.At.XFluor.Device.AMR, 2.0.10.0 Tecan.At.XFluor.Device.AMRPlus, 2.0.10.0 Tecan.At.XFluor.Device.GeniosUltra, 2.0.10.0 Tecan.At.XFluor.Device.Reader, 2.0.10.0 Tecan.At.XFluor.Device.Safire3, 2.0.10.0 Tecan.At.XFluor.Device.Safire3Pro, 2.0.10.0 Tecan.At.XFluor.Device.SunriseMini, 2.0.10.0 Tecan.At.XFluor.ExcelOutput, 2.0.10.0 Tecan.At.XFluor.NanoQuant, 2.0.10.0 Tecan.At.XFluor.ReaderEditor, 2.0.10.0
Note on E3: EHC, V_5.31_04/18_InfiniteRX (Mar 28 2018/16.41.26) MTP, V_5.31_04/18_InfiniteRX (Mar 28 2018/16.41.26) HCP, V_5.31_04/18_InfiniteRX (Mar 28 2018/16.41.26) LUM, V_2.20_02/2015_LUMINESCENCE (Feb 24 2015/16.09.02) MEM, V_3.00_09/11_MCR (Sep 27 2011/15.05.45) MEX, V_3.00_09/11_MCR (Sep 27 2011/15.05.10) ZSCAN, V_5.31_04/18_InfiniteRX (Mar 28 2018/16.41.26)

## Solution
| offsets | factor
Tip1 | 0.43857 | 1.0362
Tip2 | 0.01623 | 1.0807
Tip3 | 0.15598 | 1.0571
Tip4 | 0.084622 | 1.1277
Tip5 | 0.2182 | 1.0915
Tip6 | 0.066386 | 1.0739
Tip7 | 0.67048 | 1.0707
Tip8 | -0.42628 | 1.0933